## Slide 1
Principi fisici di radiologia

## Slide 3: Radiologia
Raggi X scoperti nel 1895 da Wilhelm Conrad Roentgen
Premio Nobel per la fisica nel 1901

## Slide 4: Cosa sono i Raggi - X ?
Radiazione elettromagnetica
Non hanno massa
Non hanno carica
Non sono visibili
Viaggiano alla velocità della luce
Possono causare fluorescenza
Possono ionizzare atomi

## Slide 5: Radiazioni elettromagnetiche
Velocità della luce costante nel vuoto
c = 3.0 x 10^8 m/s
c = λf
E = (hc) / λ

## Slide 6: Atomo
Numero di elettroni 2n^2 (K=2, L=8, M=18…)

## Slide 7: Characteristic X - Rays
Energia di un fotone trasferita ad un elettrone che viene espulso dal suo «shell»
Elettrone espulso dallo «shell» con più energia possibile

## Slide 8: Bremsstrahlung radiation
Elettrone viene deviato da un nucleo carico positivamente causando perdita di energia cinetica

## Slide 9: Produzione dei Raggi - X
Catodo –
Filamento
Anodo +
Rotante
Tungsteno
Vacuum
Rivestimento Vetro

## Slide 10: Catodo
Elettrodo negativo tubo
Filamento scalda
Produzione elettroni
Elettroni emessi dal catodo --> anodo quando applicato voltaggio

## Slide 11: Anodo
Target metallico degli elettroni
Positivo rispetto al catodo
Tungsteno (Z=74) alto punto di fusione e numero atomico
Rotante
99% calore
1% x - ray

## Slide 12: kVp e mAs
kVp è il peak di kilovoltaggio applicato al tubo
  Energia degli elettroni
mAs = mA (milliamperaggio) x tempo(s)
  Numero di elettroni
A wide spectrum of x-ray energies

## Slide 13: Anodo
99% calore
1% X-Ray
  Bremsstrahlung
  Characteristic X ray

## Slide 14: Collimazione

## Slide 15: Cosa succede se X – Ray colpisce un oggetto?
Effetto Fotoelettrico
Effetto Compton
Notes: This so-called Compton scattering happens if the energy of the incident
photon is much higher than the binding energy of the electrons of an atom.
It is the predominant effect for photons >100 kV.
This process can happen
several times within the matter, whereby the higher the physical density, the
higher the electron density and, therefore, the likelihood of interactions.
Alternatively, the photon will pass through the object without further
interaction and hit a detector system (film-screen system, CR plate, or
digital detector).

## Slide 16: Formazione della radiografia
Notes: A destra scatter, fog the fikmm e esposizione personale
Filtro rimuove x ray bassa energia

## Slide 17: Film Blackness e Opacità
Notes: A destra scatter, fog the fikmm e esposizione personale
Filtro rimuove x ray bassa energia

## Slide 18: Legge del quadrato delle distanze
Una determinata intensità è inversamente proporzionale al quadrato della distanza da quella fonte
Se la distanza da una fonte si raddoppia
L’intensità in una determinate area è ridotta ad 1/4

## Slide 19
Radiographic Film
Computed Radiography
Digital Radiography

## Slide 20: Griglie
Compton Scatter raggiungono cassetta  fog  riducono contrasto e dettaglio
mAs aumentati per compensare i raggi x primari assorbiti dalla griglia
Dispositivi per ridurre Scatter

## Slide 21: DICOM
Digital imaging and communications in medicine

## Slide 22: Radiopacità

## Slide 23: Artefatti

## Slide 24: QUANTUM MOTTLE – SOTTOESPOSIZIONE
Exposure artefact
All images created using CR or DR contain a certain amount of noise. This is mostly due to the fluctuations in the number of x-ray photons throughout the image, called quantum mottle.
The prominence of quantum mottle is dependent on the amount of data representing the object of interest in proportion to the amount of noise - referred to as the signal-to-noise ratio (SNR).
The SNR increases when the number of x-rays incident on the imaging plate or detector increases.
The use of low mAs techniques or increased attenuation of the primary x-ray beam may result in an insufficient number of incident x-rays and a low SNR.
The resultant image may have prominent quantum mottle and appear grainy, decreasing the overall image quality.
Image filtering provides a method to alter the conspicuity of quantum mottle.
  Blurring techniques reduce the prominence of quantum mottle at the expense of decreased image detail.
  Edge-enhancing techniques increase delineation of object borders as well as underlying quantum mottle.
Appropriate use of equipment-specific technique charts (to avoid underexposure) and prudent use of image filtering techniques are recommended to decrease the quantum mottle artefact.
CR or DR

## Slide 25: SATURATION - SOVRAESPOSIZIONE
Exposure artefact
DR and CR have a larger dynamic range when compared with FS.
  The lower end of the dynamic range is limited by prominence of noise, and the upper end of the dynamic range is limited by saturation.
Imaging plates and detector arrays record a latent image by storing energy proportional to the amount of x-rays incident per unit area, and there are physical limits to the amount of energy that can be stored.
When the maximum storage capacity of an F-centre (CR) or detector element (DR) has been reached, the corresponding pixel appears completely black.
Further increases in x-ray exposure have no effect on areas of saturation.
Radiograph of a canine lumbar spine acquired from a flat panel detector system. This radiograph is overexposed to the point of saturation
of the detector. The mid-abdominal organs are not seen and the soft tissues dorsal to the spine are completely black. With this degree of overexposure, the calibration mask is visible, revealing the plank-like arrangement of the flat panel detector system.
Regions of the patient with less attenuation may appear completely black before processing, and this information is not retrievable by image manipulation or enhancement techniques.
With CR systems, the saturated areas appear uniform, but some DR systems use a calibration mask to correct for nonuniformity in the detector, and with all pixels set to the same maximum value, the mask becomes visible.
  Each DR system that uses a calibration mask will typically have a characteristic pattern in the mask depending upon the construction of the panel.
Reducing exposure levels is recommended to remain within the dynamic range of the image detector and avoid saturation artefact.
CR or DR

## Slide 26: ﻿Grid Cutoff
﻿﻿﻿An incorrectly positioned focused grid will create an attenuation pattern on the image similar to that described in film-screen radiography.
﻿Correct placement of scatter reduction grids will reduce these artifacts.
CR and DR

## Slide 27: Double Expotion
Exposure/pre-exposure artefact
﻿Double Exposure
CR plates may be mistakenly used for multiple exposures or studies without being read or erased between each one. The radiograph will then appear as 2 images summated on each other. This can also occur with FS and results in double images, which are consequently overexposed. In contrast, the larger latitude of CR allows for double exposures to have a seemingly normal overall opacity. The opacity of each of the summated exposures relative to each other depends on the difference in the technique used for each one.
Double exposures have also been observed using DR as a result of electrical interruptions or data transfer error. Double exposures can be avoided by reading each imaging plate immediately after its exposure.
When using DR, scheduled maintenance and use of stable data transfer will reduce double-exposure artifact.
DR &CR

## Slide 28: FADING – Sbiadimento – quando la sviluppo dopo
Post-exposure artefact
A latent image is created using CR by exciting molecules within the imaging plate when exposed to x-rays.
The trapped energy of the excited molecules is intended to be released by a red laser during the reading process and result in a proportionate release of green light.
If an imaging plate is not read shortly after being exposed, the excited molecules in the imaging plate will gradually lose their energy and return to a neutral state.
The latent image will slowly be lost.
Reading an imaging plate several days after exposure will result in an image that may appear grainy, underexposed, or have generalized decreased quality.
Newer imaging plates have more stable forms of energy storage and are less susceptible to fading.
Fading artefacts can be avoided by immediately reading imaging plates after each exposure.
CR

## Slide 29: UBERSCHWINGER/HALO
Workstation artefact
The increased dynamic range of DR and CR is often accompanied by a decrease in image contrast (due to an increased exposure latitude). Image manipulation can be used to accentuate the margins of indistinct opacities and improve contrast. Improved delineation of object borders increases conspicuity and aids in their identification.
This is commonly performed by unsharp masking techniques. Unsharp masking is an edge enhancement algorithm that creates a border at the interface between structures of differing attenuation.
The dark border that surrounds highly attenuating objects is referred to as überschwinger, halo artefact, or rebound effect and this has been recognized in veterinary imaging.
The prominence of the surrounding dark border varies with kernel size selected during unsharp masking and the difference in attenuation between adjacent objects.
Uberschwinger is often present surrounding metallic orthopaedic implants.
It can be differentiated from pathologic changes because überschwinger is uniform in thickness, conforms to the shape of the object, and is present at other similar interfaces.
Uberschwinger becomes less pronounced with decreases in edge enhancement. Minimal use of edge enhancement is recommended when imaging body regions containing metallic implants.
CR and DR

## Slide 30: DENSITY THRESHOLD – SOGLIA DI DENSITà
Workstation artefact
Highly attenuating objects within the image are designated extremely high pixel values on the DR or CR histogram. The image is assigned a long grey scale when these extreme values are incorporated into image analysis.
  The highly attenuating object is assigned the maximum grey-scale value.
  The remaining objects in the image are near the minimum of the grey scale and do not contrast each other well.
This is referred to as a density threshold artefact and may be encountered when metallic objects are present within the image.
When processing an image, a maximum and/or minimum pixel value included in image analysis can be designated. Any object above the density threshold will appear entirely white.
Setting a maximum density threshold to exclude metallic objects will exclude their corresponding extremely high pixel values from image histogram analysis.
The image will be processed with the limited inclusion of biologic opacities, optimizing grey-scale and contrast between tissues.
The density threshold should be selected during equipment set-up
CR and DR