## Slide 1: - Food Analysis (II year) -Chemical Analysis of Foods and chemometrics (I year) a.a. 2025/26
Dario Compagnone
dcompagnone@unite.it

## Slide 2: Subjects
Mass spectrometry, principles of instrumentation, ifenation of mass spectrometry with gas and liquid chromatography. Use of LC-MS-MS and GC-MS in food analysis.
Confirmatory methods in food analysis, guidelines of Eu Legislation for food contaminants, maximum residue limits, antropic and natural contaminants, fitochemicals , control strategies, validation of methods.
IR spectroscopy, basic theory, rotational and vibrational transitions. FT-IR instrumentation, samples and sampling signal in FT-IR, qualitycontrol, qualitative and quantitative application. FT-NIR in food analysis for quality and process control.

## Slide 3: Previous knowledge required
Spectroscopy:
electromagneticwaves, intensity, amplitude, wawelenght, Beer Law, UV-VIS classical spectrophotometer
Sensors and Biosensors:
principles of potentiometry and voltammetry, enzyme kinetic (Michaelis-Menten)
Confirmatory methods and hyphenated tecniques Classical pretreatment of the samples, priciples of chromatography, gas and liquid chromatography

## Slide 4: Examination

## Slide 5: What are sensors and biosensors?
“…a biosensor is a self-contained integrated device, which is capable of providing specific quantitative or semiquantitative analytical information using a biological recognition element (biochemical receptor) which is retained in direct spatial contact with a transduction element.”
Technical report “ Reccomended definition and classification” IUPAC (Physical Chemistry and Analytical Chemistry Divisions) 2001
A sensor is a device able to transform a physical or chemical info (e.g. concentration of one or more compounds in a solution), into an analytical useful signal
Chemical sensing
Useful signal
Molecular binding
system
Biological interaction
analita
Biosensors are chemical sensors nwhich the binding element is a biological molecule

## Slide 6: Self-testing
Numbers:
85% of the global market for biosensors
70 million tests per day worldwide
15 billion $ in 2015
30 billion $ in 2024 (exp.)

## Slide 7: The biological element translates the info of the biochemical domain (e.g. concentration) into a chemical or physical signal with a certain selectivity
Transducing element
Biological binding system
selectivity
sensitivity
Tranducers
electrochemical
optical
piezoelectric
thermal
potentiometric
amperometric
voltammetric
conduttimetric, impedimetric

## Slide 8: Classification of biosensors on the basis of the biological element
Catalytic biosensors
catalysis of a chemical reaction ex: enzymes, cells, tissues
Affinity biosensors
formation of a stable complex with the ligand ex: antibodies, DNA strands, proteic receptors
Biomimetic biosensors
affinity biosensors using a synthethic receptor
es: peptides, aptamers, molularly imprinted polymers

## Slide 9: Ion selective electrodes for biosensors
The most used ISEs for the development of biosensors have
been CO2 and NH3 probes.
This potentiometric gas-sensors are realised using a pH glass electrode and a reference behind a gas permeable membrane (polytetrafluoroethilene, polypropylene, etc.). A very thin film of a suitable electrolyte is present between the gas permeable membrane and the surface of the pH electrode.
Hydrolysis of NH3 or CO2 (diffused from the sample) in the electrolyte causes a change in H+ that is measured at the electrode.
This is related to the partial pressure of the gas in the sample solution. Immobilisation of a suitable enzyme onto the surface leads to the detection of metabolites in the 10-5 – 10-2 mol/L range.

## Slide 11: Reference electrodes
calomel
Ag/AgCl

## Slide 12
Amperometric sensors monitor the current flow when a selected fixed potential is applied at a working electrode with respect to a reference electrode. The current generated by the oxidation (or reduction) of a compound is dependent by:
heterogeneous rate constant k
diffusion (mass transfer) of the electroactive specie at the
electrode surface
preceding or following chemical reactions
surface reactions (adsorption)

## Slide 14
Applied potential
-0.7 V vs. Ag/AgCl

## Slide 15
OXYGEN PROBE
The cathode should possess: high catalytic activity for the reduction of O2 , (large exchange current) , sufficient electrical conductivity (low adsorption of organic impurities or O2),it should be inert , it should exhibit a large overvoltage for the decomposition of water (no hydrogen liberation), it should permit the required construction operation (e.g. sealing into glass)
Pt and Au are the most used materials
The gas permeable membrane protects the electrodes from contamination, provides for reproducible conditions of oxygen tranport and minimize undesirable changes in electrolyte composition. Ideally, it shoud be with low permeability and high diffusivity for oxygen.
Non-hydratable polymers as polytetrafluoroethylene (PTFE), polypropylene (PP) and polyethylene (PE) are generally used. Alternatively elastic rubber or silicone are useful, even though less stable in alkaline medium.
The anode is Ag/AgCl and a neutral or alkaline electrolyte is used in the final assembling.

## Slide 16: Conventional carbon based probes:
graphite
carbon paste
pyrolitic graphite
glassy carbon
C
Front plane
Lateral plane
B
A
O
O
OH
OH

## Slide 17: Printing electrodes: serigraphy
Advantages:
Dimension
Disposable
Low-Cost
A. ink; B. squeege;
C. printing mask;
D. printing mesh;
E. frame; F. printed ink

## Slide 18: DISPOSABLE SCREEN-PRINTED CARBON ELECTRODES
counterelectrode
reference electrode
working electrode

## Slide 20
C o rre nte
P o ten ziale (V )
Cyclic voltammetries of ferricyanide (A), cathecol (B), acetamminophene (C), ascorbic acid (D) carried with dfferent screen-printed electrodes: Dupont (a), Ercon (b), Acheson
(c) Gwent (d). Supporting electrolyte; KCl 0.1M (A), scan rate 20 mV/s.

## Slide 21: Screen-Printed Electrodes

## Slide 22: How do we tune them?

## Slide 23
TISSUE BASED BIOSENSORS
bovine liver (rich in catalase)
2 H2O2  2 H2O + O2
banana (rich in polyphenol oxidase)
dopamine + O2  dopamine quinone
CELL BASED BIOSENSORS
measurement of ethanol using acetobacter xylinum (O2 electrode)
Determination of the BOD (biological oxygen demand).
The BOD values indicate the amount of biochemicallly degradable organic material (carbonaceous demand) and the oxygen used to oxidise sulphides and ferrous ion.
Conventional methods include BOD5 and BOD7 which need 5 and 7 days. BOD biosensors have been developed using Trichosporon cutaneum,
Bacillus subtilis, Hansenula anomala, etc.

## Slide 25: Spectrophotometric enzymatic kits work as end-point reaction using enzymes in solution
http://www.sigmaaldrich.com/life- science/metabolomics/enzymatic- kits.html

## Slide 26
Food Industry
Beer and Spirits
Milk and Dairy Products Fruits, Vegetables and Nuts
Meat, Poultry, Pork and Fish (Animal Protein) Wine
Drinking Water
Molecular Target Carbohydrates and Sugars Inorganic Ions
Lipids (Fats)
Amino Acids and Proteins

## Slide 27: Enzyme electrode
1
Steady state
2
time
Sample injection
S

## Slide 28
THE BIOLOGICAL ELEMENT SHOULD HAVE:
SPECIFICITY (HIGH SELECTIVITY) FOR THE ANALYTE(S)
GOOD STABILITY IN OPERATING CONDITIONS (t, pH,  )
RETENTION OF SUFFICIENT BIOLOGICAL ACTIVITY WHEN IMMOBILISED
NO (VERY LOW) INHIBITION BY THE SAMPLE

## Slide 31
E
E
E
E
E
E
E
E
E
E
E
E
E
E
E
E
E
E
E
adsorption
covalent binding
cross-linking
entrapment

## Slide 32
CHO
(CH 2 )3
CHO
O
O
O O O
O
O
N
enzima
H 2 N
O
O
O
enzima
glutaraldehyde reactions: polimerization and lysine amino group
C H 3 S O 4 -
C O C H 3
+ N H
C O 2 -
H 2 N C H ( C H 2 ) 4N H 2
C O 2 -
C N H C H ( C H 2 ) 4N
+ N H
C H
O
( C H 3 ) 3 C H
H 2 N e n z i m a
C O 2 -
C N H C H
+ N H
( C H 2 ) 4 N
H
C H ( C H 3 ) 3 C N e n z i m a
C O
N H
N Y L O N
( C H 3 ) 2 S O 4 2 -
D I M E T I L S O L F A T O
L I S I N A
C H 3 O H
C O 2 -
N H
C H ( C H 2
C
) 4N H 2
+ N H
O O
H C ( C H 3 ) 3 C H
G L U T A R A L D E I D E
C
O
OH
NH2-enzima
R' NH C O NH R''
R' N
+ C + H
N R''
R'
O NH
+
C O C
+NH R''
O H
C N enzima +
+ H +
Immobilisation via
carbodiimide

## Slide 33
The immobilised enzyme has always a higher (apparent) Km ,a shifted and larger optimum pH and lower enzymatic activity compared to the enzyme in solution

## Slide 34
Enzyme electrodes are generally classified according to the mechanism
of the electron transfer:
First generation: the enzyme is immobilised using a membrane that is in contact with the electrode surface. Usually there are other membranes to protect and regulate diffusion. Response time at the stady state is on the order of minutes
Second generation: electron transfer occurs via an electrochemical mediator in solution. The mediator shuttles electrons between the enzyme and the electrode.
Third generation: direct exchange of electrons between the electrode and the enzyme

## Slide 35
Second and third generation

## Slide 36
A good electrochemical mediator characteristics :
rapid reaction with the enzyme
rapid and reversible electron trasfer rate
low overpotial for the redox reaction
pH independent
stability in the reaction medium in both redox forms
Should not react with the dissolved oxygen in
solution
no toxicity

## Slide 37: E° = +0,53 V E° = +0,45 V E° = +0,165 V
Fe3+ + e-
Fe2+
Ferrycianide Ferrocene
ferrocene is an excellent mediator for the oxidation of glucose catalysed by glucose oxidase
glucose + GODox  gluconolactone (gluconic acid) + GODrid + 2H+
GODrid + 2Fe+  GODox 2 Fe – 2 e-  2 Fe+
+ 2Fe
Fe 3+

## Slide 40
NATURAL | E(V) vs. SHC | SYNTHETIC | E (V) vs. SHC
Cytochrome a3 | +0,29 | Esacyanoferrate(III) | +0,45
Cytochrome c3 | +0,24 | 2,6-dichlorophenol | +0,24
Ubiquinone | +0,10 | Indophenolo | +0,24
Cytochrome b | +0,08 | Ferrocene | +0,17
Vitamin K2 | -0,03 | N-metilfenazium sulphate | +0,07
Rubredoxin | -0,05 | Metilene blue | +0,4
Flavoproteins | da -0,4 a +0,2 | Ftalocyanin | -0,02
FAD/FADH2 | -0,23 | Fenosafranin | -0,23
FMN/FMNH2 | -0,23 | Benzyl viologen | -0,36
NAD+/NADH | -0,32 | Methyl viologen | -0,46
NADP+/NADPH | -0,32 |  | 
ferredossina | -0,43 |  | 

## Slide 41: Third generation Conduting salts
C
NC
CN
C
NC CN
S S
N
N
2
H C
C H 2
N H 2
N
O s
+2 / + 3
( b p y ) 2 C l
Struttura dell’ l’osmio bipiridile legato a polivinilpiridina
Tetracyanochinodimethane (TCNQ)
Redox gel
Tetra tiafulvalene(TTF)

## Slide 42: Hystamine Biosensor
PQQ Alcohol Dehydrogenase entrapped in a Os hydrogel

## Slide 43: NAD(P)H electrodes The largest class of redox enzymes known is dehydrogenases which use the NAD(P)H / NAD(P)+ couple as cofactor.
Oxidation of NADH at carbon and metal solid electrodes proceeds at high overvoltages (+400/ +700 mV vs. Ag/AgCl) via formation of the radical cation NADH.+ . This can give side reactions (dimerisation) and adsorb onto the electrode (carbon).
A soluble mediator can be used to lower the overpotential and increase the electron transfer rate
NADH + Medox  NAD+ + Medred
at the electrode surface polarised at the appropriate E :
Medred  Medox + ne
ortho- and para-quinones , quinone imines have been used and incorporated into
larger molecules as indophenols, phenazines and phenoxazines
The mediator can be also immobilized at the electrode surface giving a
chemically modified electrode for NADH
Sensitivity is excellent, major problems arise from selectivity and stability of
the sensors

## Slide 44: Reduced Prussian Blue is a selective catalyst
for H2O2 reduction
K2FeII[FeII(CN)6] + H2O2 KFeIII[FeII(CN)6] +K++ OH-
carbon electrode
Prussian Blue
oxidase enzyme
substrate
product
O2
H2O2
2H2O
2H+
2e-
Applied potential =
-0.05 V vs. Ag/AgCl

## Slide 45
-0.5
-0.25 0
0.25 0.5 0.75
E/V
0
Bare electrode
Prussian Blue modified electrode
Prussian Blue modified electrode + H2O2

## Slide 46: Selection of the applied potential
H2O2
ascorbic acid
Potential selected - 50 mV vs. Ag/AgCl pseudo-ref.
-1
0
5
4
3
2
1
-100
50
100
-50
0
i/ A
E/mV .
2%
65%
38%
20%
2%

## Slide 47: Lysine biosensor, linear range and selectivity

## Slide 48: Loop
Working buffer
Peristaltic pump
Waste
Amperometer
Recorder
Electrochemical cell
Injection valve
sample
carrier
i
t
Working electrode
Carrier solution flow
Sample injection through valve

## Slide 50: Relative activity of 3 different purified lysine oxidase
compound | Yamasa | SIGMA | Univ. of Athens
Lysine | 100 | 100 | 100
Phenylalanine | 14 | 42 | 6
Arginine | 2 | 17 | 0
Ornithine | 3 | 14 | 0
Histidine | 0 | 15 | 0
Furosine | 0 | 0 | 0
Piridosine | 0 | 0 | 0
Norleucine | 3 | 17 | 3
AGPA* | 0 | 0 | 0
Microwave hydrolysis + biosensor , analysis time 30 min
Sample | Amino-acid analysis (mM) | L-lysine biosensor (mM) | Recovery (%)
Milk | 1.684 | 1.493  0.014 | 88.66
pasta | 0.299 | 0.308  0.012 | 103.01

## Slide 51
Lactic acid monitoring during mozzarella cheese manufacturing
In mozzarella cheese manufacture L-Lactic acid is the main product of
lactose fermentation generated by selected cultures of lactic acid bacteria. A progressive acidification of the curd md the whey occurs and had to be carefully controlled.
Particularly, the pH of the "stretching point" is important in order to avoid loss of fats, a decrease in yield and low reproducibility of themanufacturing.
Optimum pH is 4.9 for water-buffalo milk and 5.1 for cow milk. At these pH values there is a great increase in the buffer capacity due to casein (isoelectric point pH ~ 5) and low molecular weight acids
A sensitive measurement of lactic acid in real time can be useful in the
optimization of the mozzarella cheese manufacture

## Slide 53: Biogenic amines
R-CH2-NH2 + O2 + H2O
R-CHO + NH3 + H2O2
0 . 2 0 . 4
[ a m m i n e ] x
0 . 6
1 0 - 4
0 . 8
m o l / L
0 . 0
1 . 0
1 . 2
corren te (nA)
0 . 4
0 . 2
0 . 0
0 . 6
1 . 2
1 . 0
0 . 8
2 . 0
1 . 8
1 . 6
1 . 4
t r i p t a m i n a
c a d a v e r i n a
i s t a m i n a t i r a m i n a
f e n i l e t a n a m m i n a
p u t r e s c i n a

## Slide 54
 |  | Storage time 20 days, temp. 0  1°C |  | 
variety | at harvest | LDPE | Super L | air
Pellecchiella |  |  |  | 
ripening time I | 5.6  0.5 | 3.0  0.2 | 2.9  0.6 | 3.4  0.6
ripening time II | 5.5  0.2 | 3.1  0.4 | 3.0  0.7 | 3.6  0.0.2
Boccuccia |  |  |  | 
ripening time I | 5.2  0.2 | 2.2  0.1 | 3.7  0.7 | 4.9  0.2
ripening time II | 5.1  0.2 | 4.6  0.8 | 4.6  0.4 | 5.1  0.1
Total amines (mg/mL), expressed as equivalents of Put, measured in apricot samples using DAO biosensora.

## Slide 55: alcohol + O2 —AOx  aldehyde + H2O2
Immobilization: PEI on Pall Immunodyne
Storage: 1% sucrose
Optimised operative conditions: 0.1 M phosphate buffer pH 7.0 + 0.02%
Tween. Flow rate 1 mL/min; injection loop 500 L.
Analitycal performances:
detection limit 10-6 mol/L linearity 2 x 10-6 / 10-3 mol/L
stability: 20% decrease after 200 samples
glycerol + ATP(Mg2+) —GK  glycerol-3-P + ADP
glycerol-3-P + O2 + H2O —GPO  glycerone-3-P + H2O2
Immobilization: GK on aminopropyl glass beads (via glutaraldehyde), GPO on
Immunodyne
Storage: DEAE-dextran/lactitol (1/5%)
Optimised operative conditions: 0.1 M borate buffer pH 8.5 + 3 mM ATP(Mg2+) + 0.02% Tween. Flow rate 0.5 mL/min; injection loop 250 L.
Analitycal performances:
detection limit 10-6 mol/L
linearity 2.5 x 10-6 / 5 x 10-4 mol/L
stability: 40% decrease after 200 samples

## Slide 56: Fructose
fructose + 2 PMS+ —FDH  ketofructose + 2 PMSH
PMSH —+100 mV (Pt)  PMS+ + e-
N
CH3SO4-
N +
CH3
Immobilization: BSA-glutaraldehyde on Immobilon AV
Storage: DEAE-dextran/lactitol (1/5%)
Optimised operative conditions: 0.1 M citrate/phosphate buffer pH 4.5 +
0.02% Tween. Flow rate 0.5 mL/min; injection loop 100 L.
Analitycal performances: detection limit 5 x 10-7 mol/L
linearity 10-6 / 8 x10-4 mol/L
stability: 30% decrease after 200 samples

## Slide 57: One delestage
glucose fructose ethanol
D
n
Ru ning off
fructose glucose ethanol
p g ver
Running off
Pumping-over

## Slide 58
glucose
fructose
ethanol
Delestage
ag
of
glucose
fructose ethanol
ag
D es
es
o

## Slide 59: Fructose:Glucose ratio during alcoholic fermentation
4
3
2
1
0
5
8
7
6
10
9
0
20
40
80
100
120
60
Time (h)
F:G
Torrevento
Solopaca
Di Majo-Norante